Title:
Unexpected air leakage due to kinking in the inflation system of Pilot Balloon Line in a Wire-Reinforced Endotracheal Tube: A Case Report

Abstract
Although rare, endotracheal tube malfunction may lead to life-threatening airway consequences during anaesthesia. We report a case of intraoperative persistent air leakage in a neurosurgical patient caused by kinking at the junction of the inflation lumen and pilot balloon line in a wire-reinforced endotracheal tube. This highlights the importance of vigilance, meticulous techniques, and thorough equipment inspection in anaesthesia practice.

Keywords: Endotracheal tube, pilot balloon, cuff malfunction, , air leak

Introduction
Airway management remains a cornerstone of anaesthesia practice. Endotracheal tubes (ETTs) are considered safe & reliable; however, equipment malfunction, although uncommon, can result in serious intraoperative complications. Malfunctions of the cuff, pilot balloon, or inflation line may mimic cuff rupture or malposition, leading to diagnostic confusion.  Frequently, the leak is not caused by a structural defect in the ETT. Cuff underinflation, cephalad migration of the ETT (partial tracheal extubation), misplaced orogastric or nasogastric tubes, wide discrepancy between ETT and tracheal diameters, or increased peak airway pressure can cause leaks around intact cuffs1. We present a case of persistent intraoperative air leakage caused by an unrecognized kink in the pilot balloon line of a reinforced ETT.

Case Report
A 42-year-old female patient was scheduled for craniotomy and excision of a glioblastoma under general anaesthesia. Following induction, tracheal intubation was performed using a 7.0 mm internal diameter wire-reinforced cuffed ETT. The initial placement was confirmed using capnography.

During fixation of the ETT, a minor delay resulted in partial displacement, causing the cuff to come little bit outwards and lie between the vocal cords. This was detected by hearing gurgling sound inside oral cavity. Decision to reposition the tube taken. The cuff was deflated and the tube advanced blindly until end-tidal CO₂ was confirmed, after which the cuff was re-inflated and the tube secured. However, significant air leakage and gurgling sounds persisted.

We suspected either cuff rupture or oesophageal intubation. As the pilot balloon remained inflated we excluded cuff rupture, but still there was persistent ETCO2 in the monitor. So we came to conclusion that tube was in trachea. We thought of inadequate cuff inflation, so we push more air in the pilot ballon but without any effect. So direct laryngoscopy done to confirm the tube position. It was found that the tube is inside trachea with no cuff visible outside glottic opening. But still air was coming out. So the ETT was removed and replaced with a new 7.5 mm cuffed wire reinforced tube, after which ventilation was adequate and no leakage occurred.

On close inspection of the removed ETT, we discovered a sharp kink at the junction of inflation lumen and the pilot balloon line, as shown in Figure 1. This prevented cuff inflation despite the pilot balloon appearing inflated.

Discussion
ETT malfunction is an under-recognised but critical cause of airway compromise in the perioperative setting. Commonly reported issues include cuff rupture, defective pilot balloon valves, or manufacturing defects in the lumen of ET tubes. Keiichi Nagasawa et.al described an incidence of endotracheal tube cuff leakage caused by a damaged pilot balloon tubing due to thermal energy of the fibre-optic light without causing any significant decrease in cuff pressure2.

In our case, a kinked inflation line at the junction of inflation lumen and pilot ballon line resulted in a unique scenario where the pilot balloon remained inflated, giving a false impression of cuff inflation.

This situation could have been catastrophic if not detected at the earliest as there was chance of aspiration and inadequate ventilation. Moreover in a neurosurgical patient with the  drapes applied it would have been almost impossible to detect at later stage. The persistence of capnography initially masked the malfunction, delaying recognition.

Key lessons include:

- Careful inspection of airway devices before use, including inflation line integrity.

- Avoiding blind advancement when malposition is suspected.

- Maintaining high suspicion of equipment malfunction when clinical findings are incongruent.

Conclusion
Even minor lapses in vigilance, compounded by rare equipment malfunction, can pose significant risk during anaesthesia. This case emphasizes the importance of meticulous airway management and thorough equipment checks to prevent catastrophic outcomes.

Financial support and sponsorship

Nil.

Conflicts of interest

There are no conflicts of interest.

Reference:

- El-Orbany, Mohammad MD; Salem, M. Ramez MD. Endotracheal Tube Cuff Leaks: Causes, Consequences, and Management. Anesthesia & Analgesia 117(2):p 428-434, August 2013. | DOI: 10.1213/ANE.0b013e318292ee21

- Nagasawa K, Nishibata M, Luthe SK, Kawamata T. A rare case of endotracheal tube cuff leakage with no detectable decrease in cuff pressure. JA Clin Rep. 2024 Nov 12;10(1):71.